## field_dup_sites_2021
monitoring_location_id | activity_type
KR RM 0 NNC | Quality Control Field Replicate Msr/Obs
KR RM 0 NNC | Field Msr/Obs
KR RM 31 | Quality Control Field Replicate Msr/Obs
KR RM 31 | Field Msr/Obs

## intrinsic_qaqc_objectives_2021
PARAMETER | METHOD_RANGE | MSR_UNITS | SENSITIVITY | SENSITIVITY_RANGE_MIN | SENSITIVITY_RANGE_MAX | PRECISION_PLUS_MINUS | ACCURACY | ACCURACY_UNITS
pH | SM H+ B-2000, Hydrolab pH probe on Minisonde 0-14 | Standard pH units | 0.01 |  |  | 0.2 | 0.2 | Standard pH units
Turbidity | ISO 7027, Self Cleaning Turbidity Hydrolab Minisonde 5 0-3000 NTU | NTU | 0.1 | 0 | 400 | 5 | 5 | percent
Turbidity | ISO 7027, Self Cleaning Turbidity Hydrolab Minisonde 5 0-3000 NTU | NTU | 1 | 400 | 3000 |  | 5 | percent
Turbidity (Soldotna Lab) | EPA 180.1 Rev 2.0, Hach 2100 P 0-1000 NTU | NTU | 0.01 | 0 | 9.99 | 5 | 5 | percent
Turbidity (Soldotna Lab) | EPA 180.1 Rev 2.0, Hach 2100 P 0-1000 NTU | NTU | 0.1 | 10 | 100 | 5 | 5 | percent
Turbidity (Soldotna Lab) | EPA 180.1 Rev 2.0, Hach 2100 P 0-1000 NTU | NTU | 1 | 100 | 1000 | 5 | 5 | percent
Water Temperature | SM 2550 B-2000 95 -5 to 95, -5 to +45, and -5 to +45 | degrees Celsius (C) | 0.1 |  |  | 0.1 | 0.1 | degrees Celsius (C)
| YSI Model 30 |  | 0.2 |  |  | 0.2 | 0.2 | degrees Celsius (C)
| YSI Model 55 |  | 0.1 |  |  | 0.2 | 0.2 | degrees Celsius (C)
Conductance | EPA 120.1, Hydrolab probe on Minesonde 0-100 mS/cm | uS/cm |  | 4 digits |  | 0.001 | 2 | percent full scale
Dissolved Oxygen | SM 4500 O C-2000, Micro Winkler Titration 0 to 20 mg/L | mg/L |  | 0.1 |  | 0.9 | 0.3 | mg/L
Dissolved Oxygen (Winkler used for calibration) | ASTM D888-09 (A), Hydrolab LDO probe on Minisonde 0 to 50 mg/L | mg/L |  | 0.01 |  | NA | 0.2 | mg/L

## analytes_qaqc_objectives_2021
PARAMETER | METHOD / RANGE | Sensitivity (MDL) | PQL | SENSITIVITY_UNIT | PROJECT_PRECISION | ACCURACY | NOTES
Gasoline Range Organics (GRO) | AK101 by GCFID: 25-1000000000 | 3 | 5x MDL |  | 20% | Calibration: Â±25% (Recovery: 60-120%)
Diesel Range Organics (DRO) | AK102 by GCFID: 50-1000000 | 0.0065 mg/L | 0.11 mg/L |  | 20% | Calibration: Â±25%
Recovery: Â±25%
Residual Range Organics (RRO) | AK103 by GCFID: 500-1000000 | 0.22 mg/L | 0.54 mg/L |  | 20% | Calibration: 75-105%
Recovery: 60-120%
Benzene | EPA 602 by GC/PID: 1-1000000000 | 0.074 Î¼g/L | 1.0 Î¼g/L |  | 20% | Calibration: Â±15%
Recovery: Â±30%
Ethylbenzene | EPA 602 by GC/PID: 1-1000000000 | 0..088 Î¼g/L | 1.0 Î¼g/L |  | 20% | Calibration: Â±15%
Recovery: Â±30%
Toluene | EPA 602 by GC/PID: 1-1000000000 | 0.078 Î¼g/L | 1.0 Î¼g/L |  | 20% | Calibration: Â±15%
Recovery: Â±30%
m,p-Xylene (total) | EPA 602 by GC/PID: 1-1000000000 | 0.20 Î¼g/L | 2.0 Î¼g/L |  | 20% | Calibration: Â±15%
Recovery: Â±30%
o-Xylene | EPA 602 by GC/PID: 1-1000000000 | 0.20 Î¼g/L | 1.0 Î¼g/L |  | 20% | Calibration: Â±15%
Recovery: Â±30%
Xylenes (total) | EPA 602 by GC/PID: 1-1000000000 | 0.82 Î¼g/L | 3.0 Î¼g/L |  | 20% | Calibration: Â±15%
Recovery: Â±30%
Bromofluorobenzene (surrogate) | EPA 602 by GCPID | 0.12 Î¼g/L | 0.50 Î¼g/L |  | 20% | Calibration: Â±15%
Recovery: Â±30%