## field_dup_sites_2021
monitoring_location_id | activity_type
KR RM 0 NNC | Quality Control Field Replicate Msr/Obs
KR RM 0 NNC | Field Msr/Obs
KR RM 31 | Quality Control Field Replicate Msr/Obs
KR RM 31 | Field Msr/Obs

## instrument_qaqc_objectives
PARAMETER | METHOD_RANGE | MSR_UNITS | SENSITIVITY | SENSITIVITY_RANGE_MIN | SENSITIVITY_RANGE_MAX | PRECISION_PLUS_MINUS | ACCURACY | ACCURACY_UNITS
pH | SM H+ B-2000, Hydrolab pH probe on Minisonde 0-14 | Standard pH units | 0.01 |  |  | 0.2 | 0.2 | Standard pH units
Turbidity | ISO 7027, Self Cleaning Turbidity Hydrolab Minisonde 5 0-3000 NTU | NTU | 0.1 | 0 | 400 | 5 | 5 | percent
Turbidity | ISO 7027, Self Cleaning Turbidity Hydrolab Minisonde 5 0-3000 NTU | NTU | 1 | 400 | 3000 |  | 5 | percent
Turbidity (Soldotna Lab) | EPA 180.1 Rev 2.0, Hach 2100 P 0-1000 NTU | NTU | 0.01 | 0 | 9.99 | 5 | 5 | percent
Turbidity (Soldotna Lab) | EPA 180.1 Rev 2.0, Hach 2100 P 0-1000 NTU | NTU | 0.1 | 10 | 100 | 5 | 5 | percent
Turbidity (Soldotna Lab) | EPA 180.1 Rev 2.0, Hach 2100 P 0-1000 NTU | NTU | 1 | 100 | 1000 | 5 | 5 | percent
Water Temperature | SM 2550 B-2000 95 -5 to 95, -5 to +45, and -5 to +45 | degrees Celsius (C) | 0.1 |  |  | 0.1 | 0.1 | degrees Celsius (C)
| YSI Model 30 |  | 0.2 |  |  | 0.2 | 0.2 | degrees Celsius (C)
| YSI Model 55 |  | 0.1 |  |  | 0.2 | 0.2 | degrees Celsius (C)
Conductance | EPA 120.1, Hydrolab probe on Minesonde 0-100 mS/cm | uS/cm |  | 4 digits |  | 0.001 | 2 | percent full scale
Dissolved Oxygen | SM 4500 O C-2000, Micro Winkler Titration 0 to 20 mg/L | mg/L |  | 0.1 |  | 0.9 | 0.3 | mg/L
Dissolved Oxygen (Winkler used for calibration) | ASTM D888-09 (A), Hydrolab LDO probe on Minisonde 0 to 50 mg/L | mg/L |  | 0.01 |  | NA | 0.2 | mg/L

## hydrocarbon_qaqc_objectives
PARAMETER | METHOD | METHOD_MIN_RANGE | METHOD_MAX_RANGE | SENSITIVITY_MDL | PQL | SENSITIVITY_UNIT | OVERALL_PROJECT_PRECISION_PCT | ACCURACY_CALIBRATION_MIN_PCT | ACCURACY_CALIBRATION_MAX_PCT | ACCURACY_RECOVERY_MIN_PCT | ACCURACY_RECOVERY_MAX_PCT | ACCURACY_PLUS_MINUS_PCT | NOTES
Gasoline Range Organics (GRO) | AK101 by GCFID | 25 | 1000000000 | 3 | 5x_MDL |  | 20% | -25% | 25% | 60% | 120% |  | PQL = 5x MDL
Diesel Range Organics (DRO) | AK102 by GCFID | 50 | 1000000 | 0.0065 | 0.11 | mg/L | 20% | -25% | 25% | -25% | 25%
Residual Range Organics (RRO) | AK103 by GCFID | 500 | 1000000 | 0.22 | 0.54 | mg/L | 20% | 75% | 105% | 60% | 120%
Benzene | EPA 602 by GC/PID | 1 | 1000000000 | 0.074 | 1 | ug/L | 20% | -15% | 15% | -30% | 30%
Ethylbenzene | EPA 602 by GC/PID | 1 | 1000000000 | 0.088 | 1 | ug/L | 20% | -15% | 15% | -30% | 30%
Toluene | EPA 602 by GC/PID | 1 | 1000000000 | 0.078 | 1 | ug/L | 20% | -15% | 15% | -30% | 30%
m,p-Xylene (total) | EPA 602 by GC/PID | 1 | 1000000000 | 0.2 | 2 | ug/L | 20% | -15% | 15% | -30% | 30%
o-Xylene | EPA 602 by GC/PID | 1 | 1000000000 | 0.2 | 1 | ug/L | 20% | -15% | 15% | -30% | 30%
Xylenes (total) | EPA 602 by GC/PID | 1 | 1000000000 | 0.82 | 3 | ug/L | 20% | -15% | 15% | -30% | 30%
Bromofluorobenzene (surrogate) | EPA 602 by GCPID |  |  | 0.12 | 0.5 | ug/L | 20% | -15% | 15% | -30% | 30%
Benzene | EPA 624 |  |  | 0.12 | 0.4 | ug/L | 20% |  |  |  |  | 20.5%
Ethylbenzene | EPA 624 |  |  | 0.31 | 1 | ug/L | 20% |  |  |  |  | 21%
Toluene | EPA 624 |  |  | 0.31 | 1 | ug/L | 20% |  |  |  |  | 20.5%
m,p-Xylene (total) | EPA 624 |  |  | 0.62 | 2 | ug/L | 20% |  |  |  |  | 20.5%
o-Xylene | EPA 624 |  |  | 0.31 | 1 | ug/L | 20% |  |  |  |  | 22%
1,2-Dichloroethane-D4 (surrogate) | EPA 624 | [10 empty] | 18.5%
4-Bromofluorobenzene (surrogate) | EPA 624 | [10 empty] | 14.4%
Toluene D-8 (surrogate) | EPA 624 | [10 empty] | 11.5%

## metals_nutrient_qaqc_objectives
PARAMETER | METHOD | SENSITIVITY_MDL | SENSITIVITY_UNIT | PQL | OVERALL_PROJECT_PRECISION_PCT | ACCURACY_PLUS_MINUS_PCT
Hardness | Calculated from Ca, Total and Mg, Total as CaCO3 + MgCO3
Calcium, Total | EPA 200.7 | 1 | mg/L | 4 | 20% | 21%
Iron, Total | EPA 200.7 | 2 | mg/L | 8 | 20% | 49.5%
Magnesium, Total | EPA 200.7 | 0.2 | mg/L | 2 | 20% | 27%
Arsenic, dissolved | EPA 200.8 | 1.5 | ug/L | 5 | 20% | 15%
Cadmium, dissolved | EPA 200.8 | 0.15 | ug/L | 0.5 | 20% | 15%
Chromium, dissolved | EPA 200.8 | 0.78 | ug/L | 2 | 20% | 15%
Copper, dissolved | EPA 200.8 | 0.031 | ug/L | 1 | 20% | 15%
Lead, dissolved | EPA 200.8 | 0.06 | ug/L | 0.2 | 20% | 15%
Zinc, dissolved | EPA 200.8 | 3.1 | ug/L | 10 | 20% | 15%
Phosphorous, Total | SM21 4500-P-B,E | 0.01 | mg/L | 0.02 | 25% | 25%
Nitrates (NO2 + NO3) | SM21 4500-NO3-F | 0.05 | mg/L | 0.2 | 25% | 10%
Total Suspended Solids | SM21-2540-+D | 0.31 | mg/L | 1 | 5% | 25%
Fecal coliform | SM 9222D-1997-membrane | 1 | cfu/100 mL | 1

## temperature_qaqc_objectives
PARAMETER | METHOD | RESOLUTION_LIMIT_C | MIN_EXPECTED_RANGE_C | MAX_EXPECTED_RANGE_C | ACCURACY_PLUS_MINUS_C | PRECISION_C | COMPLETENESS_PCT | NOTES
AIR TEMPERATURE | HOBO TIDBIT V2 | 0.16 | -5 | 37 | 0.2 | 0.4 | 90% | Other method: SM 2550 B-2000
WATER TEMPERATURE | HOBO WATER TEMP PRO V2 | 0.02 | 0 | 50 | 0.2 | 0.4 | 90% | Accuracy within 0-50 C range